## Slide 1: Force and strain measurement

## Slide 2: Force and strain measurement methods
Resistance strain gauges
  Metal foil gauges
  Semiconductor gauges
  Thin films
Load cells
  Strain gauge based
  Piezoelectric based
LVDTs, Extensometers for displacement and/or strain measurement
Capacitive gauges
Moirè strain measurement
Digital Image Correlation
Interferometry

## Slide 3: Strain gage principles

## Slide 4
Take the total differential of the resistance and normalize by the resistance.

## Slide 5: Using Wheatstone bridge to measure resistance changes
R1
R2
R3
R4
Ei
Eo

## Slide 6: Quarter bridge

## Slide 7: Half bridge

## Slide 8: Half bridge v.2

## Slide 9: ¾ Bridge
No such thing!

## Slide 10: Full bridge

## Slide 11: Using half and full bridges to increase signal and to cancel out thermal strains
F

## Slide 12: Using half to increase signal and to cancel out bending strains

## Slide 14: Load cells use full bridges to remove lead wire and thermal strain issues

## Slide 16: Reading load cell output
Differential amplifier having a gain, G = RF/R1.
(R1, R2, and R3 are not the bridge resistors).
Load cells are usually described in terms of output for full scale load. For instance, if the calibration sheet for a FS = 5000 N full scale load cell states the sensitivity is SLC = 3.2 mV/V, this means that the output of the bridge will be 3.2 mV for a volt of excitation (Eex).
The output of the bridge is always amplified with a differential amplifier because the signal is small and because the voltage is measured between two points neither of which is ground.
The output voltage for a given force is:
The force for a measured voltage is:

## Slide 17: Piezoelectric effect – crystal deformation displaces the relative center of charge resulting in the development of charge across the faces.

## Slide 19: Charge amplifier used to convert charge to voltage by integrating the current over a specified time. It will not detect a static force.

## Slide 20
Typical maximum frequency for a strain gauge load cell is 2-8kHz Hz. Piezoelectric load cells range from 0.3 to 100,000 Hz.
Full scale deflection on the order of 0.1 mm for strain gauge load cells. PZT load cells are typically an order of magnitude stiffer.

## Slide 22: LVDT – linear variable differential transformer Measures displacements with high displacement resolution.

## Slide 23: Moire methods

## Slide 25
= 6 degrees + 2.5% strain
 = 6 degrees

## Slide 26
p’=0.47 in
-10 degree rotation
p=0.4 in

## Slide 27: Out of plane displacement measurement

## Slide 28: Digital image correlation

## Slide 29: Example from Correlated Solutions web site

## Slide 30: LIMESS web site has a few videos http://www.limess.com/index.php/en/company/technology/digital-image-correlation http://www.limess.com/index.php/en/applications/deformation-strain-measurement
Left image from http://trilion.com/when-is-a-digital-image-correlation-system-the-right-tool/
Right image from http://lib.znate.ru/docs/index-99125.html

## Slide 33: Particle imaging velocimetry

## Slide 35: Electronic speckle pattern interferometry

## Slide 36: Laser speckle – arises from random interference between all scattering points on surface
Assume the phase at every point is independent of the other points. This suggests the average U will follow Gaussian statistics. The irradiance probability is described by where the most likely probability is an intensity of zero. This changes when two beams are added together

## Slide 37: Out-of-plane displacement measurement using speckle interferometry

## Slide 40: Photoelasticity
Stress-optic law

## Slide 42: Photoelastic coatings